## COMPONENTS
|  | COMPONENT | LENGTH (mm) | CLEAR interface connector | PHYTRON 1:
ANALYZER | Y TRANSLATION | PHYTRON 1 | 2550 | MAYT, EAYT | ZSS32,200,1,2+HD08/100
|  | RENISHAW 1 | 3670 |  | SIZE 32, RESOLUTION 200
| X ROTATION | PHYTRON 1 | 3920 | MAXR, EAXR | I12-CURRENT 1,2 Amp
|  | RENISHAW 3 | 4200 |  | HARMONIC GEAR i=100
| Z TRANSLATION | PHYTRON 1 | 4200 | MAZT, EAZT
|  | RENISHAW 1 | 4900 |  | HUBER 1:
| BENDERS | PRECISTEP 1 | 5450 | MEAB1 | MOTOR PHYTRON ZSS43.500-2.5:
|  | LIMIT SWITCH 2 | 5450 |  | 14 Ncm, flange 42 mm, 500/1000 steps/rev.
|  | LIMIT SWITCH 2 | 5450 |  | incl. wiring and connector.
|  | PRECISTEP 1 | 5450 | MEAB2
|  | LIMIT SWITCH 2 | 5450 |  | HUBER 2:
|  | LIMIT SWITCH 2 | 5450 |  | MOTOR PHYTRON ZSS43.500-2.5:
DETECTOR | Z TRANSLATION | PHYTRON 1 | 1940 | MDZT | 14 Ncm, flange 42 mm, 500/1000 steps/rev.
|  | LIMIT SWITCH 1 | 1940 |  | incl. wiring and connector.
|  | LIMIT SWITCH 1 | 1940 |  | 2 limit switches (PA 6.6) with Al-housing.
|  | RENISHAW 1 |  |  | ZERO POINT 9100
| Y TRANSLATION | PHYTRON 1 | 3510 | MDYT
|  | LIMIT SWITCH 1 | 3510 |  | PRECISTEP 1:
|  | LIMIT SWITCH 1 | 3510 |  | AM2224-2R-AV-18-14+22/5-69.2:1
|  | RENISHAW 1 |  |  | with encoder: PE 22-120
| X TRANSLATION | PRECISTEP 2 | 4930 | MDX1T
|  | LIMIT SWITCH 1 | 4930 |  | PRECISTEP 2:
|  | LIMIT SWITCH 1 | 4930 |  | AM2224-2R-AV-18-14+22/5-69.2:1
| X TRANSLATION | PRECISTEP 2 | 5250 | MDX2T
|  | LIMIT SWITCH 1 | 5250 |  | PIN:
|  | LIMIT SWITCH 1 | 5960 |  | HAMAMATSU
| X ROTATION | HUBER 1 | 5250 | MDXR | S3590-08
|  | RENISHAW 2 | 5250 | EDXR
|  | PIN | 5610 |  | DETECTOR:
|  | DETECTOR | 5660 |  | DECTRIS
SLITS | X OPEN | PRECISTEP 2 | 2870 | MOTX1 | MYTHEN 1K
|  | LIMIT SWITCH 1 | 2870
|  | LIMIT SWITCH 1 | 2870 |  | RENISHAW 1:
|  | PRECISTEP 2 | 2620 | MOTX2 | TONiC™ RGSZ linear encoder system:
|  | LIMIT SWITCH 1 | 2620 |  | Readhead: T1000-30A
|  | LIMIT SWITCH 1 | 2620 |  | Scale: RGSZ20u-xxx
| Z OPEN | PRECISTEP 2 | 2440 | MOTZ1 | Interface: Ti-0020-A20-A
|  | LIMIT SWITCH 1 | 2440
|  | LIMIT SWITCH 1 | 2440 |  | RENISHAW 2:
|  | PRECISTEP 2 | 2340 | MOTZ2 | TONiC™ RGSZ rotary (angle) encoder system
|  | LIMIT SWITCH 1 | 2340 |  | Readhead: T2011-30A
|  | LIMIT SWITCH 1 | 2340 |  | Scale: RGSZ20u-300mm
| X ROTATION | HUBER 2 | 2860 | MOTXR | Interface: Ti-0040-A20-A
RENISHAW 3:
SiGNUM™ RESM rotary (angle) encoder system
Readhead: SR030A
Scale: RESM20USA150
Interface: Si-NN-0200
LIMIT SWITCH 1:
BAUMER IFRM 03N3501/L
LIMIT SWITCH 2:
OMRON SS 5

## MOTOR CONFIG
MOTOR | Reference | Current (A) | Resistance | Poles | Phases | Steps/turn | Reduction | Steps/axis_turn
Phytron 1 | ZSS32,200,1,2+HD08/100 | 1.2 | 2.6 | 50 | 2 | 200 | 100 | 20000
Huber 1 | ZSS43.500-2.5 | 2.5 | 5.2 | 125 | 2 | 500 | 1 | 500
Huber 2 | ZSS43.500-2.5 | 2.5 | 5.2 | 125 | 2 | 500 | 1 | 500
Precistep 1 | AM2224-2R-AV-18-14+22/5-69.2:3-PE22-120 | 0.25 | 18 | 6 | 2 | 24 | 69.2 | 1660.8
Precistep 2 | AM2224-2R-AV-18-14+22/5-69.2:3 | 0.25 | 18 | 6 | 2 | 24 | 69.2 | 1660.8

## AXIS CALC
| MOTOR |  |  |  |  | AXIS
AXIS | TYPE | FULL_STEP |  | MICRO STEP defined |  | SPINDLE |  | Movement Resolution (m) | ENCODER SYSTEM |  |  |  |  | SW unit (m) | Steps per unit | LIM + |  | LIM -
|  | steps/turn_motor | steps/turn_axis | steps/turn_motor | steps/turn_axis | TYPE | STEP (m) |  | READHEAD | INTERFACE | RESOLUTION | encoder_counts/turn_axis | encoder_counts/turn_motor |  |  | Indexer | Encoder | Indexer | Encoder
AYT | Phytron1 | 200 | 20000 | 400 | 40000 | BNK1404 | 0.004 | 1e-07 | T1000-30A | Ti-0020-A20-A | 1e-06 | 4000 | 40 | 0.001 | 10000 | 1005580 | 100561 | -3268897 | -326854
AZT | Phytron1 | 200 | 20000 | 400 | 40000 | BNK1404 | 0.004 | 1e-07 | T1000-30A | Ti-0020-A20-A | 1e-06 | 4000 | 40 | 0.001 | 10000 | 3337495 | 333754 | -1159076 | -115908
DZT | Phytron1 | 200 | 20000 | 400 | 40000 | BNK1404 | 0.004 | 1e-07 | T1000-30A | Ti-0020-A20-A | 1e-06 | 4000 | 40 | 0.001 | 10000 | 2963613 | 296341 | -1489870 | -148974
DYT | Phytron1 | 200 | 20000 | 400 | 40000 | BNK1404 | 0.004 | 1e-07 | T1000-30A | Ti-0020-A20-A | 1e-06 | 4000 | 40 | 0.001 | 10000 | 1185178 | 118518 | -6020250 | -602025
DX1T | Precistep2 | 24 | 1660.8 | 48 | 3321.6 |  | 0.001 | 3.0106e-07 | [5 empty] | 0.001 | 3321.6 | 135733 |  | 0
DX2T | Precistep2 | 24 | 1660.8 | 48 | 3321.6 |  | 0.001 | 3.0106e-07 | [5 empty] | 0.001 | 3321.6 | 134488 |  | 0
SX1 | Precistep2 | 24 | 1660.8 | 48 | 3321.6 |  | 0.001 | 3.0106e-07 | [5 empty] | 0.001 | 3321.6 | 169390 |  | 0
SX2 | Precistep2 | 24 | 1660.8 | 48 | 3321.6 |  | 0.001 | 3.0106e-07 | [5 empty] | 0.001 | 3321.6 | 0 |  | -167100
SZ1 | Precistep2 | 24 | 1660.8 | 48 | 3321.6 |  | 0.001 | 3.0106e-07 | [5 empty] | 0.001 | 3321.6 | 167376 |  | 0
SZ2 | Precistep2 | 24 | 1660.8 | 48 | 3321.6 |  | 0.001 | 3.0106e-07 | [5 empty] | 0.001 | 3321.6 | 0 |  | -169636
AB1
AB2
The SW unit is chosen to be 1mm
| MOTOR |  |  |  |  | AXIS
AXIS | TYPE | FULL_STEP |  | MICRO STEP defined |  | Rotation |  | Movement Resolution (º) | ENCODER SYSTEM |  |  |  |  | SW unit (m) | Steps per unit | LIM + |  | LIM -
|  | steps/turn_motor | steps/º | steps/turn_motor | steps/º | R (m) | Arc in 1º (m) |  | READHEAD | INTERFACE | RESOLUTION | encoder_counts/º | encoder_counts/turn_motor |  |  | Indexer | Encoder | Indexer | Encoder
DXR | Huber1 | 500 | 250 | 2000 | 1000 | 0.055 | 0.00095993 | 0.001 | T2011-30A | Ti-0040-A20-A | 5e-07 | 1919.8622 | 3839.7244 | 1 | 1000 | 91699 | 176252 | 0 | 0
SXR | Huber2 | 500 | 250 | 2000 | 1000 | 0.055 | 0.00095993 | 0.001 | [5 empty] | 1 | 1000 | 50261 |  | 0
| MOTOR |  |  |  |  | LINEAL AXIS |  |  | ANGULAR AXIS
AXIS | TYPE | FULL_STEP |  | MICRO STEP defined |  | SPINDLE |  | Movement Resolution (m) | Rotation |  | Movement Resolution (º) | ENCODER SYSTEM | [5 empty] | SW unit (º) | Steps per unit | LIM + |  | LIM -
|  | steps/turn_motor | steps/turn_axis | steps/turn_motor | steps/turn_axis | TYPE | STEP (m) |  | R (m) | Arc in 1º (m) |  | READHEAD | INTERFACE | RESOLUTION | encoder_counts/turn_axis | encoder_counts/turn_motor | encoder_counts/º |  |  | Indexer | Encoder | Indexer | Encoder
AXR | Phytron1 | 200 | 20000 | 400 | 40000 | BNK1402 | 0.002 | 5e-08 | 0.075 | 0.001309 | Not lineal | SR030-A | Si-NN-0200 | 1e-07 | Not lineal. Average is 11300*40000/50000= 9040 | Not lineal. Average is 9040/100=90.4 | 13089.9694 | 1 | Not lineal. Depends on the position of the motor. See Indexer-Angle curve in sheet AXR | 2980990 | 672923 | -50420 | -9894
See AXR sheet.

## AYT
AYT
PICTURES
AXIS, MOTOR and ENCODER configuration
Encoder configuration based on calculations done in AXIS CALC tab. As the motor is driven in half_stepping (400 steps/turn), with a motion velocity of 4000 stps/s, the motor is moving at 10 turns/s. This means 400 encoder_steps/s, which is an encoder velocity low enough to be sure that no steps are not lost.
TESTS
TEST #1
A homing sequence is performed with a speed velocity of 1000 steps/s. When the homing signal is found in positive direction, the axis position (encoder) is set to 0. Then the axis is moved up to the lim+ position, and the Indexer and Encoder registers are recorded. From this point, the axis is moved in the full range (from lim+ to lim-) and the registers are recorded. This full range movement is repeated many times in order to see if motor or encoder pulses are lost. The motion velocity of this test is 4000 steps/s
| Move | LIM | Register | Value | Trip
| 1 | lim+ | Indexer | 1005632
|  |  | Encoder | 100567
| 2 | lim- | Indexer | -3268879 | 4274511
|  |  | Encoder | -326851 | 427418 | um
| 3 | lim+ | Indexer | 1005580 | 4274459
|  |  | Encoder | 100561 | 427412 | um
| 4 | lim- | Indexer | -3268897 | 4274477
|  |  | Encoder | -326854 | 427415 | um
With these results it can be observed that the maximum deviation in the encoder counts in the full range is 6 counts (6um). It can also be observed that to do the full range, the indexer pulses needed vary up to in 52 pulses. Besides, the relation between Indexer and encoder registers should be ideally 10, and it isn't. It makes sense that some indexer pulses are lost in the operation. Feasability of operating in closed loop should be studied.
TEST #2
After performing TEST 1, the axis is moved to encoder position -100000 and the indexer is set to -1000000. In this point the axis is moved several times with +/-500000 indexer pulses. After every movement the encoder and indexer registers are registered. The motion velocity used in this test is 4000 steps/s and 8000 steps/s, and the acceleration time is 0.8s and 1.6s. Through all the combinations it can be determined if the movement errors are produced by the velocity, acceleration time or mechanical non-linearitites.
| Move | Register | Value | Trip | @ 4000 steps/s, 0.8s acc. Time |  |  |  |  | Move | Register | Value | Trip | @ 4000 steps/s, 1.6s acc. Time
| 0 | Indexer | -1000000 | [6 empty] | 0 | Indexer | -1000000
|  | Encoder | -100001 | [7 empty] | Encoder | -100000
| 1 | Indexer | -500000 | 500000 | [5 empty] | 1 | Indexer | -500000 | 500000
|  | Encoder | -50012 | 49989 | [6 empty] | Encoder | -50016 | 49984
| 2 | Indexer | -1000000 | 500000 | [5 empty] | 2 | Indexer | -1000000 | 500000
|  | Encoder | -99997 | 49985 | [6 empty] | Encoder | -100000 | 49984
| 3 | Indexer | -1500000 | 500000 | [5 empty] | 3 | Indexer | -1500000 | 500000
|  | Encoder | -149994 | 49997 | [6 empty] | Encoder | -149998 | 49998
| 4 | Indexer | -1000000 | 500000 | [5 empty] | 4 | Indexer | -1000000 | 500000
|  | Encoder | -100006 | 49988 | [6 empty] | Encoder | -100009 | 49989
| 5 | Indexer | -500000 | 500000 | [5 empty] | 5 | Indexer | -500000 | 500000
|  | Encoder | -50012 | 49994 | [6 empty] | Encoder | -50015 | 49994
| 6 | Indexer | -1000000 | 500000 | [5 empty] | 6 | Indexer | -1000000 | 500000
|  | Encoder | -99997 | 49985 | [6 empty] | Encoder | -100000 | 49985
| Move | Register | Value | Trip | @ 8000 steps/s, 0.8s acc. Time |  |  |  |  | Move | Register | Value | Trip | @ 8000 steps/s, 1.6s acc. Time
| 0 | Indexer | -1000000 | [6 empty] | 0 | Indexer | -1000000
|  | Encoder | -100000 | [7 empty] | Encoder | -100000
| 1 | Indexer | -500000 | 500000 | [5 empty] | 1 | Indexer | -500000 | 500000
|  | Encoder | -50016 | 49984 | [6 empty] | Encoder | -50015 | 49985
| 2 | Indexer | -1000000 | 500000 | [5 empty] | 2 | Indexer | -1000000 | 500000
|  | Encoder | -100000 | 49984 | [6 empty] | Encoder | -99999 | 49984
| 3 | Indexer | -1500000 | 500000 | [5 empty] | 3 | Indexer | -1500000 | 500000
|  | Encoder | -149998 | 49998 | [6 empty] | Encoder | -149997 | 49998
| 4 | Indexer | -1000000 | 500000 | [5 empty] | 4 | Indexer | -1000000 | 500000
|  | Encoder | -100009 | 49989 | [6 empty] | Encoder | -100009 | 49988
| 5 | Indexer | -500000 | 500000 | [5 empty] | 5 | Indexer | -500000 | 500000
|  | Encoder | -50015 | 49994 | [6 empty] | Encoder | -50015 | 49994
| 6 | Indexer | -1000000 | 500000 | [5 empty] | 6 | Indexer | -1000000 | 500000
|  | Encoder | -100000 | 49985 | [6 empty] | Encoder | -99998 | 49983
Comparison
[5 empty] | Move 0 | Trip | Move 1 | Trip | Move 2 | Trip | Move 3 | Trip | Move 4 | Trip | Move 5 | Trip | Move 6
| Case 1. 4000 steps/s, 0.8s acc.time |  |  | Indexer | -1000000 | 500000 | -500000 | -500000 | -1000000 | -500000 | -1500000 | 500000 | -1000000 | 500000 | -500000 | -500000 | -1000000
|  |  |  | Encoder | -100001 | 49989 | -50012 | -49985 | -99997 | -49997 | -149994 | 49988 | -100006 | 49994 | -50012 | -49985 | -99997
| Case 2. 4000 steps/s, 1.6s acc.time |  |  | Indexer | -1000000 | 500000 | -500000 | -500000 | -1000000 | -500000 | -1500000 | 500000 | -1000000 | 500000 | -500000 | -500000 | -1000000
|  |  |  | Encoder | -100000 | 49984 | -50016 | -49984 | -100000 | -49998 | -149998 | 49989 | -100009 | 49994 | -50015 | -49985 | -100000
| Case 3. 8000 steps/s, 0.8s acc.time |  |  | Indexer | -1000000 | 500000 | -500000 | -500000 | -1000000 | -500000 | -1500000 | 500000 | -1000000 | 500000 | -500000 | -500000 | -1000000
|  |  |  | Encoder | -100000 | 49984 | -50016 | -49984 | -100000 | -49998 | -149998 | 49989 | -100009 | 49994 | -50015 | -49985 | -100000
| Case 4. 8000 steps/s, 1.6s acc.time |  |  | Indexer | -1000000 | 500000 | -500000 | -500000 | -1000000 | -500000 | -1500000 | 500000 | -1000000 | 500000 | -500000 | -500000 | -1000000
|  |  |  | Encoder | -100000 | 49985 | -50015 | -49984 | -99999 | -49998 | -149997 | 49988 | -100009 | 49994 | -50015 | -49983 | -99998
Through the comparison of the results, it can be seen that the movement errors (diference between indexer pulses and encoder pulses -should be exactly 10 times less that indexer pulses-) are not affected by varying the motion velocity or acceleration time. As the movement errors are repeatable in the axis range when changing the velocity, shows that maybe the source of errors is a mechanical issue or mechanical non-linearity. The proposed way to control the axis avoiding as much as posible the movement errors, is to implement the closed-loop control. It is decided to use the 8000 steps/s velocity and 0.8s of acc.time.
TEST #3
The motor is configured in closed-loop with the following configuration:
As the position source is configured by the TGTENC multiplying the encoder value by the relation between motor_steps/turn and encoder_steps/turn, we have: Position=EncIn*400/40=EncIn*10
So the resolution of the Deadband and Settle window must be 10. Both parameters are set to 10 in order to have 1um accuracy.
With this configuration, the axis is moved setting the target position from a position near the lim+ (1005000) up to a position near lim- (-3268000). The movement is repeated 4 times, and it is checked that the closed loop works correctly and that the encoder and indexer registers match. So this motor configuration in closed loop is considered to be OK.
AYT
lim -
lim +

## AZT
AZT
PICTURES
AXIS, MOTOR, ENCODER and CLOSED LOOP configuration
The applied configuration is the one agreed to be OK after all the testing of AYT axis.
Encoder configuration based on calculations done in AXIS CALC tab. As the motor is driven in half_stepping (400 steps/turn), with a motion velocity of 8000 stps/s, the motor is moving at 20 turns/s. This means 800 encoder_steps/s, which is an encoder velocity low enough to be sure that no steps are not lost.
TESTS
TEST #1
A homing sequence is performed with a speed velocity of 500 steps/s (in open loop). When the homing signal is found in positive direction, the axis position (encoder) is set to 0. Then the axis is moved up to the lim+ position, and the Indexer and Encoder registers are recorded. From this point, the axis is moved in the full range (from lim+ to lim-) and the registers are recorded. The motion velocity of this test is 8000 steps/s
| Move | LIM | Register | Value | Trip
| 1 | lim+ | Indexer | 3337495
|  |  | Encoder | 333754
| 2 | lim- | Indexer | -1159076 | 4496571
|  |  | Encoder | -115908 | 449662 | um
TEST #2
The axis has been previously configured in closed-loop.
With this configuration, the axis is moved setting the target position from a position near the lim+ (3330000) up to a position near lim- (-1150000). The movement is repeated twice, and it is checked that the closed loop works correctly and that the encoder and indexer registers match. So this motor configuration in closed loop is considered to be OK.
AZT
lim -
lim +

## DYT
DYT
PICTURES
AXIS, MOTOR, ENCODER and CLOSED LOOP configuration
The applied configuration is the one agreed to be OK after all the testing of AYT axis.
Encoder configuration based on calculations done in AXIS CALC tab. As the motor is driven in half_stepping (400 steps/turn), with a motion velocity of 8000 stps/s, the motor is moving at 20 turns/s. This means 800 encoder_steps/s, which is an encoder velocity low enough to be sure that no steps are not lost.
TESTS
TEST #1
A homing sequence is performed with a speed velocity of 500 steps/s (in open loop). When the homing signal is found in positive direction, the axis position (encoder) is set to 0. Then the axis is moved up to the lim+ position, and the Indexer and Encoder registers are recorded. From this point, the axis is moved in the full range (from lim+ to lim-) and the registers are recorded. The motion velocity of this test is 8000 steps/s
| Move | LIM | Register | Value | Trip
| 1 | lim+ | Indexer | 1185178
|  |  | Encoder | 118518
| 2 | lim- | Indexer | -6020250 | 7205428
|  |  | Encoder | -602025 | 720543 | um
TEST #2
The axis has been previously configured in closed-loop.
With this configuration, the axis is moved setting the target position from a position near the lim+ (1170000) up to a position near lim- (-6000000). The movement is repeated twice, and it is checked that the closed loop works correctly and that the encoder and indexer registers match. So this motor configuration in closed loop is considered to be OK.
lim -
lim +
DYT

## DZT
DZT
PICTURES
AXIS, MOTOR, ENCODER and CLOSED LOOP configuration
The applied configuration is the one agreed to be OK after all the testing of AYT axis.
Encoder configuration based on calculations done in AXIS CALC tab. As the motor is driven in half_stepping (400 steps/turn), with a motion velocity of 8000 stps/s, the motor is moving at 20 turns/s. This means 800 encoder_steps/s, which is an encoder velocity low enough to be sure that no steps are not lost.
TESTS
TEST #1
A homing sequence is performed with a speed velocity of 500 steps/s (in open loop). When the homing signal is found in positive direction, the axis position (encoder) is set to 0. Then the axis is moved up to the lim+ position, and the Indexer and Encoder registers are recorded. From this point, the axis is moved in the full range (from lim+ to lim-) and the registers are recorded. The motion velocity of this test is 8000 steps/s
| Move | LIM | Register | Value | Trip
| 1 | lim+ | Indexer | 2963613
|  |  | Encoder | 296341
| 2 | lim- | Indexer | -1489870 | 4453483
|  |  | Encoder | -148974 | 445315 | um
TEST #2
The axis has been previously configured in closed-loop.
With this configuration, the axis is moved setting the target position from a position near the lim+ (2950000) up to a position near lim- (-1470000). The movement is repeated twice, and it is checked that the closed loop works correctly and that the encoder and indexer registers match. So this motor configuration in closed loop is considered to be OK.
lim -
lim +
DZT

## AXR
AXR
PICTURES
AXIS, MOTOR, ENCODER and CLOSED LOOP configuration
The initial configuration for this axis is:
TESTS
TEST #1
A homing sequence is performed with a speed velocity of 250 steps/s (in open loop). When the homing signal is found in positive direction, the axis position (encoder) is set to 0. Then the axis is moved up to the lim+ position, and the Indexer and Encoder registers are recorded. From this point, the axis is moved in the full range (from lim+ to lim-) and the registers are recorded. The motion velocity of this test is 4000 steps/s
| Move | LIM | Register | Value | Trip
| 1 | lim+ | Indexer | 2980990
|  |  | Encoder | 672923
| 2 | lim- | Indexer | -50420 | 3031410
|  |  | Encoder | -9894 | 682817
TEST #2
In the previous test it was observed that the encoder counts do not progress linearly with the indexer. So, a closed loop configuration is very difficult to implement unless the indexer-encoder non-linear curve is well characterized.
To do so, the escript motor_characterization done my Marc Rosanes is applied to this axis. The characterization of the motor is done forcing a indexer value (the axis is moved) and then recording the encoder value.
In this case, the characterization is done from indexer value -50000 up to 2950000, with steps of 50000. (pyhton motor_characterization icesat01 16 -50000 2950000 50000). Previously the driver is configured to have the position source as indexer.
The results are shown following:
| Indexer values: | Encoder values: | Encoder Increment | Encoder Error per increment | Accumulated Encoder Error
| -50000 | -9821
| 0 | 6 | 9827 | -1473.7797 | -1473.7797
| 50000 | 9969 | 9963 | -1337.7797 | -2811.5593
| 100000 | 20054 | 10085 | -1215.7797 | -4027.339
| 150000 | 30275 | 10221 | -1079.7797 | -5107.1186
| 200000 | 40625 | 10350 | -950.7797 | -6057.8983
| 250000 | 51084 | 10459 | -841.7797 | -6899.678
| 300000 | 61637 | 10553 | -747.7797 | -7647.4576
| 350000 | 72298 | 10661 | -639.7797 | -8287.2373
| 400000 | 83077 | 10779 | -521.7797 | -8809.0169
| 450000 | 93970 | 10893 | -407.7797 | -9216.7966
| 500000 | 104983 | 11013 | -287.7797 | -9504.5763
| 550000 | 116091 | 11108 | -192.7797 | -9697.3559
| 600000 | 127273 | 11182 | -118.7797 | -9816.1356
| 650000 | 138542 | 11269 | -31.7797 | -9847.9153
| 700000 | 149900 | 11358 | 57.2203 | -9790.6949
| 750000 | 161347 | 11447 | 146.2203 | -9644.4746
| 800000 | 172877 | 11530 | 229.2203 | -9415.2542
| 850000 | 184494 | 11617 | 316.2203 | -9099.0339
| 900000 | 196182 | 11688 | 387.2203 | -8711.8136
| 950000 | 207935 | 11753 | 452.2203 | -8259.5932
| 1000000 | 219750 | 11815 | 514.2203 | -7745.3729
| 1050000 | 231621 | 11871 | 570.2203 | -7175.1525
| 1100000 | 243541 | 11920 | 619.2203 | -6555.9322
| 1150000 | 255504 | 11963 | 662.2203 | -5893.7119
| 1200000 | 267505 | 12001 | 700.2203 | -5193.4915
| 1250000 | 279544 | 12039 | 738.2203 | -4455.2712
| 1300000 | 291612 | 12068 | 767.2203 | -3688.0508
| 1350000 | 303693 | 12081 | 780.2203 | -2907.8305
| 1400000 | 315793 | 12100 | 799.2203 | -2108.6102
| 1450000 | 327891 | 12098 | 797.2203 | -1311.3898
| 1500000 | 340003 | 12112 | 811.2203 | -500.1695
| 1550000 | 352097 | 12094 | 793.2203 | 293.0508
| 1600000 | 364190 | 12093 | 792.2203 | 1085.2712
| 1650000 | 376265 | 12075 | 774.2203 | 1859.4915
| 1700000 | 388319 | 12054 | 753.2203 | 2612.7119
| 1750000 | 400340 | 12021 | 720.2203 | 3332.9322
| 1800000 | 412325 | 11985 | 684.2203 | 4017.1525
| 1850000 | 424266 | 11941 | 640.2203 | 4657.3729
| 1900000 | 436156 | 11890 | 589.2203 | 5246.5932
| 1950000 | 448002 | 11846 | 545.2203 | 5791.8136
| 2000000 | 459781 | 11779 | 478.2203 | 6270.0339
| 2050000 | 471510 | 11729 | 428.2203 | 6698.2542
| 2100000 | 483159 | 11649 | 348.2203 | 7046.4746
| 2150000 | 494735 | 11576 | 275.2203 | 7321.6949
| 2200000 | 506243 | 11508 | 207.2203 | 7528.9153
| 2250000 | 517671 | 11428 | 127.2203 | 7656.1356
| 2300000 | 529023 | 11352 | 51.2203 | 7707.3559
| 2350000 | 540271 | 11248 | -52.7797 | 7654.5763
| 2400000 | 551417 | 11146 | -154.7797 | 7499.7966
| 2450000 | 562454 | 11037 | -263.7797 | 7236.0169
| 2500000 | 573376 | 10922 | -378.7797 | 6857.2373
| 2550000 | 584220 | 10844 | -456.7797 | 6400.4576
| 2600000 | 594956 | 10736 | -564.7797 | 5835.678
| 2650000 | 605590 | 10634 | -666.7797 | 5168.8983
| 2700000 | 616110 | 10520 | -780.7797 | 4388.1186
| 2750000 | 626486 | 10376 | -924.7797 | 3463.339
| 2800000 | 636732 | 10246 | -1054.7797 | 2408.5593
| 2850000 | 646884 | 10152 | -1148.7797 | 1259.7797
| 2900000 | 656925 | 10041 | -1259.7797 | -4.0018e-11
| Mean of the encoder increment for a indexer increment of 50000 steps: | [6 empty] | 11300.7797 | encoder counts
| 50000 | indexer steps are |  | 125 | motor_turns
| So, there are: |  | 4.4245 | indexer_steps/encoder_counts
| So there are : |  | 90.4062 | encoder_counts/motor_turn
TEST #3
The following closed loop configuration was tested:
The maximum error steps in the closed loop were configured to be 10000 (maximum accumulated error obtained in the previous test).
With this configuration the movements are imprecise and the position after the movement does not match with the target position.
If the maximum error steps parameter is set to 100 (usual value) the closed loop stops the movment because the instantaneous error exceed the threshold.
As a conclusion, we can obtain that this axis can not be driven in closed loop. It will be driven in open loop, after homing calibration.
The SW control will have to operate in open_loop, forcing the indexer value to reach a certain encoder position (which will mean a certain angle).
So, final configuration will be:
TEST #4
A final Indexer - Encoder characterization is performend, more precisely than in Test 2. In this case the script is executed from indexer position -45000 up to 2975000 with 151 steps of 20000.
With the information from AXIS CALC sheet, where it is calculated that there are 13090 encoder_counts/º, the curve Indexer-angle is calculated.
Indexer values: | Encoder values: | Encoder Increment | Encoder Error per increment | Accumulated Encoder Error | Angle value | Angle Increment (º) | Angle Error per increment | Accumulated Angle Error
-50420 | -9894 |  |  |  | -0.75584
-45000 | -8867 |  |  |  | -0.67739
-25000 | -4947 | 3920 | -591.14 | -591.14 | -0.37792 | 0.29947 | -0.04516 | -0.04516
-5000 | -996 | 3951 | -560.14 | -1151.28 | -0.076089 | 0.30183 | -0.042791 | -0.087951 |  | Mean of the encoder increment for a indexer increment of 20000 steps: | 4511.14 | encoder counts
15000 | 2971 | 3967 | -544.14 | -1695.42 | 0.22697 | 0.30306 | -0.041569 | -0.084361
35000 | 6952 | 3981 | -530.14 | -2225.56 | 0.53109 | 0.30413 | -0.0405 | -0.082069 |  | Mean of the angle increment for a indexer increment of 20000 steps: | 0.34462
55000 | 10950 | 3998 | -513.14 | -2738.7 | 0.83652 | 0.30542 | -0.039201 | -0.079701
75000 | 14977 | 4027 | -484.14 | -3222.84 | 1.1442 | 0.30764 | -0.036985 | -0.076186
95000 | 19020 | 4043 | -468.14 | -3690.98 | 1.453 | 0.30886 | -0.035763 | -0.072749
115000 | 23090 | 4070 | -441.14 | -4132.12 | 1.7639 | 0.31092 | -0.033701 | -0.069464
135000 | 27178 | 4088 | -423.14 | -4555.26 | 2.0762 | 0.3123 | -0.032325 | -0.066026
155000 | 31284 | 4106 | -405.14 | -4960.4 | 2.3899 | 0.31367 | -0.03095 | -0.063276
175000 | 35417 | 4133 | -378.14 | -5338.54 | 2.7057 | 0.31574 | -0.028888 | -0.059838
195000 | 39569 | 4152 | -359.14 | -5697.68 | 3.0228 | 0.31719 | -0.027436 | -0.056324
215000 | 43732 | 4163 | -348.14 | -6045.82 | 3.3409 | 0.31803 | -0.026596 | -0.054032
235000 | 47917 | 4185 | -326.14 | -6371.96 | 3.6606 | 0.31971 | -0.024915 | -0.051511
255000 | 52112 | 4195 | -316.14 | -6688.1 | 3.9811 | 0.32047 | -0.024151 | -0.049066
275000 | 56321 | 4209 | -302.14 | -6990.24 | 4.3026 | 0.32154 | -0.023082 | -0.047233
295000 | 60550 | 4229 | -282.14 | -7272.38 | 4.6257 | 0.32307 | -0.021554 | -0.044636
315000 | 64799 | 4249 | -262.14 | -7534.52 | 4.9503 | 0.3246 | -0.020026 | -0.04158
335000 | 69061 | 4262 | -249.14 | -7783.66 | 5.2759 | 0.32559 | -0.019033 | -0.039059
355000 | 73346 | 4285 | -226.14 | -8009.8 | 5.6032 | 0.32735 | -0.017276 | -0.036309
375000 | 77646 | 4300 | -211.14 | -8220.94 | 5.9317 | 0.3285 | -0.01613 | -0.033406
395000 | 81964 | 4318 | -193.14 | -8414.08 | 6.2616 | 0.32987 | -0.014755 | -0.030885
415000 | 86301 | 4337 | -174.14 | -8588.22 | 6.5929 | 0.33132 | -0.013303 | -0.028058
435000 | 90651 | 4350 | -161.14 | -8749.36 | 6.9252 | 0.33231 | -0.01231 | -0.025613
455000 | 95027 | 4376 | -135.14 | -8884.5 | 7.2595 | 0.3343 | -0.010324 | -0.022634
475000 | 99432 | 4405 | -106.14 | -8990.64 | 7.596 | 0.33652 | -0.0081085 | -0.018432
495000 | 103842 | 4410 | -101.14 | -9091.78 | 7.9329 | 0.3369 | -0.0077265 | -0.015835
515000 | 108273 | 4431 | -80.14 | -9171.92 | 8.2714 | 0.3385 | -0.0061222 | -0.013849
535000 | 112719 | 4446 | -65.14 | -9237.06 | 8.6111 | 0.33965 | -0.0049763 | -0.011099
555000 | 117174 | 4455 | -56.14 | -9293.2 | 8.9514 | 0.34034 | -0.0042888 | -0.0092651
575000 | 121639 | 4465 | -46.14 | -9339.34 | 9.2925 | 0.3411 | -0.0035248 | -0.0078136
595000 | 126108 | 4469 | -42.14 | -9381.48 | 9.6339 | 0.34141 | -0.0032193 | -0.0067441
615000 | 130618 | 4510 | -1.14 | -9382.62 | 9.9785 | 0.34454 | -8.7089e-05 | -0.0033063
635000 | 135136 | 4518 | 6.86 | -9375.76 | 10.3236 | 0.34515 | 0.00052406 | 0.00043697
655000 | 139666 | 4530 | 18.86 | -9356.9 | 10.6697 | 0.34607 | 0.0014408 | 0.0019649
675000 | 144209 | 4543 | 31.86 | -9325.04 | 11.0167 | 0.34706 | 0.0024339 | 0.0038747
695000 | 148760 | 4551 | 39.86 | -9285.18 | 11.3644 | 0.34767 | 0.0030451 | 0.005479
715000 | 153317 | 4557 | 45.86 | -9239.32 | 11.7125 | 0.34813 | 0.0035034 | 0.0065485
735000 | 157918 | 4601 | 89.86 | -9149.46 | 12.064 | 0.35149 | 0.0068648 | 0.010368
755000 | 162515 | 4597 | 85.86 | -9063.6 | 12.4152 | 0.35118 | 0.0065592 | 0.013424
775000 | 167127 | 4612 | 100.86 | -8962.74 | 12.7675 | 0.35233 | 0.0077051 | 0.014264
795000 | 171740 | 4613 | 101.86 | -8860.88 | 13.1199 | 0.35241 | 0.0077815 | 0.015487
815000 | 176374 | 4634 | 122.86 | -8738.02 | 13.4739 | 0.35401 | 0.0093858 | 0.017167
835000 | 181014 | 4640 | 128.86 | -8609.16 | 13.8284 | 0.35447 | 0.0098442 | 0.01923
855000 | 185678 | 4664 | 152.86 | -8456.3 | 14.1847 | 0.3563 | 0.011678 | 0.021522
875000 | 190344 | 4666 | 154.86 | -8301.44 | 14.5412 | 0.35646 | 0.01183 | 0.023508
895000 | 195021 | 4677 | 165.86 | -8135.58 | 14.8985 | 0.3573 | 0.012671 | 0.024501
915000 | 199712 | 4691 | 179.86 | -7955.72 | 15.2568 | 0.35837 | 0.01374 | 0.026411
935000 | 204407 | 4695 | 183.86 | -7771.86 | 15.6155 | 0.35867 | 0.014046 | 0.027786
955000 | 209115 | 4708 | 196.86 | -7575 | 15.9752 | 0.35966 | 0.015039 | 0.029085
975000 | 213840 | 4725 | 213.86 | -7361.14 | 16.3361 | 0.36096 | 0.016338 | 0.031377
995000 | 218564 | 4724 | 212.86 | -7148.28 | 16.697 | 0.36089 | 0.016261 | 0.032599
1015000 | 223297 | 4733 | 221.86 | -6926.42 | 17.0586 | 0.36157 | 0.016949 | 0.03321
1035000 | 228042 | 4745 | 233.86 | -6692.56 | 17.4211 | 0.36249 | 0.017866 | 0.034814
1055000 | 232796 | 4754 | 242.86 | -6449.7 | 17.7843 | 0.36318 | 0.018553 | 0.036419
1075000 | 237564 | 4768 | 256.86 | -6192.84 | 18.1485 | 0.36425 | 0.019623 | 0.038176
1095000 | 242336 | 4772 | 260.86 | -5931.98 | 18.5131 | 0.36455 | 0.019928 | 0.039551
1115000 | 247111 | 4775 | 263.86 | -5668.12 | 18.8778 | 0.36478 | 0.020157 | 0.040086
1135000 | 251892 | 4781 | 269.86 | -5398.26 | 19.2431 | 0.36524 | 0.020616 | 0.040773
1155000 | 256680 | 4788 | 276.86 | -5121.4 | 19.6089 | 0.36578 | 0.02115 | 0.041766
1175000 | 261488 | 4808 | 296.86 | -4824.54 | 19.9762 | 0.3673 | 0.022678 | 0.043829
1195000 | 266291 | 4803 | 291.86 | -4532.68 | 20.3431 | 0.36692 | 0.022296 | 0.044975
1215000 | 271106 | 4815 | 303.86 | -4228.82 | 20.7109 | 0.36784 | 0.023213 | 0.04551
1235000 | 275918 | 4812 | 300.86 | -3927.96 | 21.0785 | 0.36761 | 0.022984 | 0.046197
1255000 | 280729 | 4811 | 299.86 | -3628.1 | 21.4461 | 0.36753 | 0.022908 | 0.045892
1275000 | 285554 | 4825 | 313.86 | -3314.24 | 21.8147 | 0.3686 | 0.023977 | 0.046885
1295000 | 290385 | 4831 | 319.86 | -2994.38 | 22.1837 | 0.36906 | 0.024435 | 0.048413
1315000 | 295217 | 4832 | 320.86 | -2673.52 | 22.5529 | 0.36914 | 0.024512 | 0.048947
1335000 | 300053 | 4836 | 324.86 | -2348.66 | 22.9223 | 0.36944 | 0.024817 | 0.049329
1355000 | 304878 | 4825 | 313.86 | -2034.8 | 23.2909 | 0.3686 | 0.023977 | 0.048794
1375000 | 309720 | 4842 | 330.86 | -1703.94 | 23.6608 | 0.3699 | 0.025276 | 0.049253
1395000 | 314556 | 4836 | 324.86 | -1379.08 | 24.0303 | 0.36944 | 0.024817 | 0.050093
1415000 | 319407 | 4851 | 339.86 | -1039.22 | 24.4008 | 0.37059 | 0.025963 | 0.050781
1435000 | 324245 | 4838 | 326.86 | -712.36 | 24.7704 | 0.3696 | 0.02497 | 0.050934
1455000 | 329089 | 4844 | 332.86 | -379.5 | 25.1405 | 0.37005 | 0.025429 | 0.050399
1475000 | 333928 | 4839 | 327.86 | -51.64 | 25.5102 | 0.36967 | 0.025047 | 0.050475
1495000 | 338771 | 4843 | 331.86 | 280.22 | 25.8801 | 0.36998 | 0.025352 | 0.050399
1515000 | 343611 | 4840 | 328.86 | 609.08 | 26.2499 | 0.36975 | 0.025123 | 0.050475
1535000 | 348455 | 4844 | 332.86 | 941.94 | 26.6199 | 0.37005 | 0.025429 | 0.050552
1555000 | 353288 | 4833 | 321.86 | 1263.8 | 26.9892 | 0.36921 | 0.024588 | 0.050017
1575000 | 358126 | 4838 | 326.86 | 1590.66 | 27.3587 | 0.3696 | 0.02497 | 0.049558
1595000 | 362959 | 4833 | 321.86 | 1912.52 | 27.728 | 0.36921 | 0.024588 | 0.049558
1615000 | 367794 | 4835 | 323.86 | 2236.38 | 28.0973 | 0.36937 | 0.024741 | 0.049329
1635000 | 372624 | 4830 | 318.86 | 2555.24 | 28.4663 | 0.36898 | 0.024359 | 0.0491
1655000 | 377454 | 4830 | 318.86 | 2874.1 | 28.8353 | 0.36898 | 0.024359 | 0.048718
1675000 | 382275 | 4821 | 309.86 | 3183.96 | 29.2036 | 0.3683 | 0.023672 | 0.048031
1695000 | 387093 | 4818 | 306.86 | 3490.82 | 29.5717 | 0.36807 | 0.023442 | 0.047114
1715000 | 391902 | 4809 | 297.86 | 3788.68 | 29.939 | 0.36738 | 0.022755 | 0.046197
1735000 | 396716 | 4814 | 302.86 | 4091.54 | 30.3068 | 0.36776 | 0.023137 | 0.045892
1755000 | 401517 | 4801 | 289.86 | 4381.4 | 30.6736 | 0.36677 | 0.022144 | 0.04528
1775000 | 406316 | 4799 | 287.86 | 4669.26 | 31.0402 | 0.36662 | 0.021991 | 0.044134
1795000 | 411101 | 4785 | 273.86 | 4943.12 | 31.4057 | 0.36555 | 0.020921 | 0.042912
1815000 | 415886 | 4785 | 273.86 | 5216.98 | 31.7713 | 0.36555 | 0.020921 | 0.041843
1835000 | 420660 | 4774 | 262.86 | 5479.84 | 32.136 | 0.36471 | 0.020081 | 0.041002
1855000 | 425445 | 4785 | 273.86 | 5753.7 | 32.5015 | 0.36555 | 0.020921 | 0.041002
1875000 | 430208 | 4763 | 251.86 | 6005.56 | 32.8654 | 0.36387 | 0.019241 | 0.040162
1895000 | 434966 | 4758 | 246.86 | 6252.42 | 33.2289 | 0.36348 | 0.018859 | 0.038099
1915000 | 439707 | 4741 | 229.86 | 6482.28 | 33.5911 | 0.36218 | 0.01756 | 0.036419
1935000 | 444450 | 4743 | 231.86 | 6714.14 | 33.9534 | 0.36234 | 0.017713 | 0.035273
1955000 | 449179 | 4729 | 217.86 | 6932 | 34.3147 | 0.36127 | 0.016643 | 0.034356
1975000 | 453908 | 4729 | 217.86 | 7149.86 | 34.6759 | 0.36127 | 0.016643 | 0.033286
1995000 | 458620 | 4712 | 200.86 | 7350.72 | 35.0359 | 0.35997 | 0.015345 | 0.031988
2015000 | 463317 | 4697 | 185.86 | 7536.58 | 35.3947 | 0.35882 | 0.014199 | 0.029543
2035000 | 468007 | 4690 | 178.86 | 7715.44 | 35.753 | 0.35829 | 0.013664 | 0.027862
2055000 | 472688 | 4681 | 169.86 | 7885.3 | 36.1106 | 0.3576 | 0.012976 | 0.02664
2075000 | 477357 | 4669 | 157.86 | 8043.16 | 36.4673 | 0.35668 | 0.01206 | 0.025036
2095000 | 482021 | 4664 | 152.86 | 8196.02 | 36.8236 | 0.3563 | 0.011678 | 0.023737
2115000 | 486657 | 4636 | 124.86 | 8320.88 | 37.1778 | 0.35416 | 0.0095386 | 0.021216
2135000 | 491291 | 4634 | 122.86 | 8443.74 | 37.5318 | 0.35401 | 0.0093858 | 0.018924
2155000 | 495901 | 4610 | 98.86 | 8542.6 | 37.884 | 0.35218 | 0.0075523 | 0.016938
2175000 | 500526 | 4625 | 113.86 | 8656.46 | 38.2373 | 0.35332 | 0.0086982 | 0.016251
2195000 | 505124 | 4598 | 86.86 | 8743.32 | 38.5885 | 0.35126 | 0.0066356 | 0.015334
2215000 | 509701 | 4577 | 65.86 | 8809.18 | 38.9382 | 0.34966 | 0.0050313 | 0.011667
2235000 | 514266 | 4565 | 53.86 | 8863.04 | 39.2869 | 0.34874 | 0.0041146 | 0.0091459
2255000 | 518819 | 4553 | 41.86 | 8904.9 | 39.6348 | 0.34782 | 0.0031979 | 0.0073125
2275000 | 523348 | 4529 | 17.86 | 8922.76 | 39.9807 | 0.34599 | 0.0013644 | 0.0045623
2295000 | 527886 | 4538 | 26.86 | 8949.62 | 40.3274 | 0.34668 | 0.0020519 | 0.0034163
2315000 | 532395 | 4509 | -2.14 | 8947.48 | 40.6719 | 0.34446 | -0.00016348 | 0.0018885
2335000 | 536887 | 4492 | -19.14 | 8928.34 | 41.015 | 0.34316 | -0.0014622 | -0.0016257
2355000 | 541366 | 4479 | -32.14 | 8896.2 | 41.3572 | 0.34217 | -0.0024553 | -0.0039175
2375000 | 545829 | 4463 | -48.14 | 8848.06 | 41.6982 | 0.34095 | -0.0036776 | -0.0061329
2395000 | 550276 | 4447 | -64.14 | 8783.92 | 42.0379 | 0.33972 | -0.0048999 | -0.0085775
2415000 | 554714 | 4438 | -73.14 | 8710.78 | 42.3769 | 0.33904 | -0.0055875 | -0.010487
2435000 | 559124 | 4410 | -101.14 | 8609.64 | 42.7138 | 0.3369 | -0.0077265 | -0.013314
2455000 | 563513 | 4389 | -122.14 | 8487.5 | 43.0491 | 0.33529 | -0.0093308 | -0.017057
2475000 | 567888 | 4375 | -136.14 | 8351.36 | 43.3833 | 0.33422 | -0.0104 | -0.019731
2495000 | 572246 | 4358 | -153.14 | 8198.22 | 43.7163 | 0.33293 | -0.011699 | -0.022099
2515000 | 576601 | 4355 | -156.14 | 8042.08 | 44.049 | 0.3327 | -0.011928 | -0.023627
2535000 | 580943 | 4342 | -169.14 | 7872.94 | 44.3807 | 0.3317 | -0.012921 | -0.02485
2555000 | 585262 | 4319 | -192.14 | 7680.8 | 44.7106 | 0.32995 | -0.014678 | -0.0276
2575000 | 589576 | 4314 | -197.14 | 7483.66 | 45.0402 | 0.32956 | -0.01506 | -0.029739
2595000 | 593857 | 4281 | -230.14 | 7253.52 | 45.3672 | 0.32704 | -0.017581 | -0.032642
2615000 | 598129 | 4272 | -239.14 | 7014.38 | 45.6936 | 0.32636 | -0.018269 | -0.03585
2635000 | 602391 | 4262 | -249.14 | 6765.24 | 46.0192 | 0.32559 | -0.019033 | -0.037302
2655000 | 606630 | 4239 | -272.14 | 6493.1 | 46.343 | 0.32383 | -0.02079 | -0.039823
2675000 | 610845 | 4215 | -296.14 | 6196.96 | 46.665 | 0.322 | -0.022623 | -0.043413
2695000 | 615038 | 4193 | -318.14 | 5878.82 | 46.9853 | 0.32032 | -0.024304 | -0.046927
2715000 | 619211 | 4173 | -338.14 | 5540.68 | 47.3041 | 0.31879 | -0.025832 | -0.050136
2735000 | 623363 | 4152 | -359.14 | 5181.54 | 47.6213 | 0.31719 | -0.027436 | -0.053268
2755000 | 627492 | 4129 | -382.14 | 4799.4 | 47.9367 | 0.31543 | -0.029193 | -0.056629
2775000 | 631604 | 4112 | -399.14 | 4400.26 | 48.2509 | 0.31413 | -0.030492 | -0.059685
2795000 | 635696 | 4092 | -419.14 | 3981.12 | 48.5635 | 0.31261 | -0.03202 | -0.062512
2815000 | 639770 | 4074 | -437.14 | 3543.98 | 48.8747 | 0.31123 | -0.033395 | -0.065415
2835000 | 643829 | 4059 | -452.14 | 3091.84 | 49.1848 | 0.31008 | -0.034541 | -0.067936
2855000 | 647873 | 4044 | -467.14 | 2624.7 | 49.4937 | 0.30894 | -0.035687 | -0.070228
2875000 | 651898 | 4025 | -486.14 | 2138.56 | 49.8012 | 0.30749 | -0.037138 | -0.072825
2895000 | 655905 | 4007 | -504.14 | 1634.42 | 50.1073 | 0.30611 | -0.038513 | -0.075652
2915000 | 659896 | 3991 | -520.14 | 1114.28 | 50.4122 | 0.30489 | -0.039736 | -0.078249
2935000 | 663859 | 3963 | -548.14 | 566.14 | 50.715 | 0.30275 | -0.041875 | -0.08161
2955000 | 667804 | 3945 | -566.14 | -6.3665e-12 | 51.0163 | 0.30138 | -0.04325 | -0.085125
2980990 | 672923 |  |  |  | 51.4074
lim -
lim +
AXR

## DXR
DXR
PICTURES
AXIS, MOTOR and ENCODER configuration
Encoder configuration based on calculations done in AXIS CALC tab.
TESTS
TEST #1
This axis has no homing, so the homing is assigned to lim- signal. Then the axis is moved up to the lim- position, and the Indexer and Encoder registers are set to 0. From this point, the axis is moved in the full range (from lim- to lim+) and the registers are recorded. The motion velocity of this test is 1000 steps/s (0.5 motor_turn/s = 1º/s).
The used motor in this axis is a Huber 409 X3 W1, which moves 2º/motor_turn. In this case the motor is driven with a micro_stepping of 2000 steps/motor_turn, which means a resolution 0.001º/step.
| Move | LIM | Register | Value | Trip
| 1 | lim- | Indexer | 0
|  |  | Encoder | 0
| 2 | lim+ | Indexer | 91699 | 91699 | 45.8495 | motor_turns
|  |  | Encoder | 176252 | 176252 | 3844.1422 | encoder_counts/motor_turn
As shown above, the encoder_counts/motor_turn parameter is slighlty different from the calculated theoretically. Maybe, it is because the encoder rule height was not taken into account in the calulations. The encoder_counts/motor_turn parameter is finally set to 3844.
TEST #2
With the corrected encoder definition, the driver is configured is closed_loop as following:
With this configuration, the axis is moved setting the target position from a position near the lim+ (85000) up to a position near lim- (1000). The movement is repeated 4 times, and it is checked that the closed loop works correctly and that the encoder and indexer registers match. So this motor configuration in closed loop is considered to be OK.
lim -
lim +
DXR

## SXR
SXR
PICTURES
AXIS and MOTOR configuration
THIS AXIS HAS NO ENCODER
TESTS
TEST #1
This axis has no homing, so the homing is assigned to lim- signal. Then the axis is moved up to the lim- position, and the Indexer and Encoder registers are set to 0. From this point, the axis is moved in the full range (from lim- to lim+) and the registers are recorded. The motion velocity of this test is 1000 steps/s (0.5 motor_turn/s = 1º/s).
The used motor in this axis is a Huber 409 X3 W1.hv, which moves 2º/motor_turn. In this case the motor is driven with a micro_stepping of 2000 steps/motor_turn, which means a resolution 0.001º/step.
| Move | LIM | Register | Value | Trip
| 1 | lim- | Indexer | 0
| 2 | lim+ | Indexer | 50261 | 50261 | 25.1305 | motor_turns
lim -
lim +

## DX1T DX2T
DX1T DX2T
PICTURES
AXIS, MOTOR and HOMING configuration
DX1T | [8 empty] | DX2T
Encoder and Closed_loop configurations are not shown because are irrelevant. No encoder is used with this motor. | [8 empty] | Encoder and Closed_loop configurations are not shown because are irrelevant. No encoder is used with this motor.
TESTS
TEST #1
The homing sequence is performed manually in both axis, moving the motor in negative direction until Lim- is reached. Then, the indexer register is set to 0. After that, the motor is moved in positive direction until Lim+ is reached.
| DX1T |  |  |  | DX2T
| LIM | Register | Value | Trip | LIM | Register | Value | Trip
| lim+ | Indexer | 135733 |  | lim+ | Indexer | 134488
| lim- | Indexer | 0 | 135733 | lim- | Indexer | 0 | 134488
DX1T
DX2T
lim +
lim -
lim +
lim -

## SX1 SX2
SX1 SX2
PICTURES
AXIS, MOTOR and HOMING configuration
SX1 | [8 empty] | SX2
Encoder and Closed_loop configurations are not shown because are irrelevant. No encoder is used with this motor. | [8 empty] | Encoder and Closed_loop configurations are not shown because are irrelevant. No encoder is used with this motor.
TESTS
TEST #1
The homing sequence is performed manually in both axis, moving the motor in negative direction until Lim- is reached (case of SX1) or moving the motor in positive direction until Lim+ is reached (case of SX2). Then, the indexer register is set to 0 for both drivers. After that, the motor is moved in positive direction until Lim+ is reached (case of SX1) or moved in negative direction until Lim- is reached (case of SX2).
With this configuration of homing for both axis, the home position is in totally open position.
| SX1 |  |  |  | SX2
| LIM | Register | Value | Trip | LIM | Register | Value | Trip
| lim+ | Indexer | 169390 |  | lim+ | Indexer | 0
| lim- | Indexer | 0 | 169390 | lim- | Indexer | -167100 | 167100
lim +
lim -
SX1
SX2
lim +
lim -

## SZ1 SZ2
SZ1 SZ2
PICTURES
AXIS, MOTOR and HOMING configuration
SZ1 | [8 empty] | SZ2
Encoder and Closed_loop configurations are not shown because are irrelevant. No encoder is used with this motor. | [8 empty] | Encoder and Closed_loop configurations are not shown because are irrelevant. No encoder is used with this motor.
TESTS
TEST #1
The homing sequence is performed manually in both axis, moving the motor in negative direction until Lim- is reached (case of SZ1) or moving the motor in positive direction until Lim+ is reached (case of SZ2). Then, the indexer register is set to 0 for both drivers. After that, the motor is moved in positive direction until Lim+ is reached (case of SZ1) or moved in negative direction until Lim- is reached (case of SZ2).
With this configuration of homing for both axis, the home position is in totally open position.
| SZ1 |  |  |  | SZ2
| LIM | Register | Value | Trip | LIM | Register | Value | Trip
| lim+ | Indexer | 167376 |  | lim+ | Indexer | 0
| lim- | Indexer | 0 | 167376 | lim- | Indexer | -169636 | 169636
lim +
lim -
SX1
SX2
lim +
lim -

## ANALYZER SIGNALS
| DB37-1 | ANMOT |  | DB37-2 |  | ANENC |  | PHYTRON |  | PRECISTEP_1 |  | PRECISTEP_2 |  | LIMIT_SWITCH_1 |  | LIMIT_SWITCH_2 |  | RENISHAW_EAYT | RENISHAW_EAZT | RENISHAW_EAXR
1 | MAYT.PHA+ | MAYT.PHA+ | 1 | 5V | 1 | 5V | 1 | MAYT.PHA+ | 1 | MEAB1.PHA+ | 1 | MEAB2.PHA+ | 1 | LPAB1 | 1 | LPAB2 | 1 | NC | NC | NC
2 | MAYT.PHA- | MAYT.PHA- | 2 | EAYT.A+ | 2 | EAYT.A- | 2 | MAYT.PHA- | 2 | MEAB1.PHA- | 2 | MEAB2.PHA- | 2 | 0V | 2 | 0V | 2 | 0V | 0V | 0V
3 | MAYT.PHA+ | MAYT.PHB+ | 3 | EAYT.A- | 3 | EAYT.B- | 3 | MAYT.PHB+ | 3 | MEAB1.PHB+ | 3 | MEAB2.PHB+ | 3 | LNAB1 | 3 | LNAB2 | 3 | NC | NC | NC
4 | MAYT.PHA- | MAYT.PHB- | 4 | EAYT.B+ | 4 | EAYT.Z- | 4 | MAYT.PHB- | 4 | MEAB1.PHB- | 4 | MEAB2.PHB- | 4 | 0V | 4 | 0V | 4 | EAYT.Z- | EAZT.Z- | EAXR.Z-
5 | MAYT.PHB+ | MAXR.PHA+ | 5 | EAYT.B- | 5 | EAYT.Q | 5 | MAXR.PHA+ | 5 | 5V | 5 | 5V |  |  |  |  | 5 | EAYT.B- | EAZT.B- | EAXR.B-
6 | MAYT.PHB- | MAXR.PHA- | 6 | EAYT.Z+ | 6 | 5V | 6 | MAXR.PHA- | 6 | 0V | 6 | 0V |  |  |  |  | 6 | EAYT.A- | EAZT.A- | EAXR.A-
7 | MAYT.PHB+ | MAXR.PHB+ | 7 | EAYT.Z- | 7 | EAZT.A- | 7 | MAXR.PHB+ | 7 | MEAB1.EA | 7 | MEAB2.EA |  |  |  |  | 7 | 5V | 5V | 5V
8 | MAYT.PHB- | MAXR.PHB- | 8 | EAYT.P | 8 | EAZT.B- | 8 | MAXR.PHB- | 8 | MEAB1.EB | 8 | MEAB2.EB |  |  |  |  | 8 | 5V | 5V | 5V
9 | MAXR.PHA+ | MAZT.PHA+ | 9 | EAYT.Q | 9 | EAZT.Z- | 9 | MAZT.PHA+ | 9 | NC | 9 | NC |  |  |  |  | 9 | 0V | 0V | 0V
10 | MAXR.PHA- | MAZT.PHA- | 10 | 0V | 10 | EAZT.Q | 10 | MAZT.PHA- | 10 | NC | 10 | NC |  |  |  |  | 10 | EAYT.Q | EAZT.Q | EAXR.Q
11 | MAXR.PHA+ | MAZT.PHB+ | 11 | 5V | 11 | 5V | 11 | MAZT.PHB+ | [8 empty] | 11 | EAYT.P | EAZT.P | EAXR.P
12 | MAXR.PHA- | MAZT.PHB- | 12 | EAXR.A+ | 12 | EAXR.A- | 12 | MAZT.PHB- | [8 empty] | 12 | EAYT.Z+ | EAZT.Z+ | EAXR.Z+
13 | MAXR.PHB+ | MEAB1.PHA+ | 13 | EAXR.A- | 13 | EAXR.B- | [10 empty] | 13 | EAYT.B+ | EAZT.B+ | EAXR.B+
14 | MAXR.PHB- | MEAB1.PHA- | 14 | EAXR.B+ | 14 | EAXR.Z- | [10 empty] | 14 | EAYT.A+ | EAZT.A+ | EAXR.A+
15 | MAXR.PHB+ | MEAB1.PHB+ | 15 | EAXR.B- | 15 | EAXR.Q | [10 empty] | 15 | NC | NC | NC
16 | MAXR.PHB- | MEAB1.PHB- | 16 | EAXR.Z+ | 16 | 5V
17 | MAZT.PHA+ | 0V | 17 | EAXR.Z- | 17 | MEAB1.EB
18 | MAZT.PHA- | LPAB1 | 18 | EAXR.P | 18 | MEAB2.EA
19 | MAZT.PHA+ | LNAB1 | 19 | EAXR.Q | 19 | 0V
20 | MAZT.PHA- | MAYT.PHA+ | 20 | 0V | 20 | EAYT.A+
21 | MAZT.PHB+ | MAYT.PHA- | 21 | 5V | 21 | EAYT.B+
22 | MAZT.PHB- | MAYT.PHB+ | 22 | EAZT.A+ | 22 | EAYT.Z+
23 | MAZT.PHB+ | MAYT.PHB- | 23 | EAZT.A- | 23 | EAYT.P
24 | MAZT.PHB- | MAXR.PHA+ | 24 | EAZT.B+ | 24 | 0V
25 | MEAB1.PHA+ | MAXR.PHA- | 25 | EAZT.B- | 25 | EAZT.A+
26 | MEAB1.PHA- | MAXR.PHB+ | 26 | EAZT.Z+ | 26 | EAZT.B+
27 | MEAB1.PHB+ | MAXR.PHB- | 27 | EAZT.Z- | 27 | EAZT.Z+
28 | MEAB1.PHB- | MAZT.PHA+ | 28 | EAZT.P | 28 | EAZT.P
29 | MEAB2.PHA+ | MAZT.PHA- | 29 | EAZT.Q | 29 | 0V
30 | MEAB2.PHA- | MAZT.PHB+ | 30 | 0V | 30 | EAXR.A+
31 | MEAB2.PHB+ | MAZT.PHB- | 31 | MEAB1.EA | 31 | EAXR.B+
32 | MEAB2.PHB- | MEAB2.PHA+ | 32 | MEAB1.EB | 32 | EAXR.Z+
33 | 0V | MEAB2.PHA- | 33 | 0V | 33 | EAXR.P
34 | LPAB1 | MEAB2.PHB+ | 34 | MEAB2.EA | 34 | 0V
35 | LNAB1 | MEAB2.PHB- | 35 | MEAB2.EB | 35 | MEAB1.EA
36 | LPAB2 | LPAB2 | 36 | 0V | 36 | 0V
37 | LNAB2 | LNAB2 | 37 | 5V | 37 | MEAB2.EB

## DETECTOR SIGNALS
| DETMOT |  | DETMOT(2) |  | DETENC |  | DETENC(2) |  | PHYTRON |  | PRECISTEP |  | LIMIT_SWITCH_1 |  | LIMIT_SWITCH_2 |  | HUBER |  | RENISHAW |  | DET DB9 | DET DB50 |  | RENISHAW_DYT | RENISHAW_DZT
1 | MDZT.PHA+ | 1 | MDZT.PHA+ | 1 | 24V | 1 | 24V | 1 | MDZT.PHA+ | 1 | MDX1.PHA+ | 1 | 24V | 1 | 24V | 1 | MDXR.PHA- | 1 | NC | 1 | PIN.+12VL | DET (50) | 1 | NC | NC
2 | MDZT.PHA- | 2 | MDZT.PHA- | 2 | LPDZT.OUT | 2 | LNDZT.OUT | 2 | MDZT.PHA- | 2 | MDX1.PHA- | 2 | LPDZT.OUT | 2 | LPDX1.OUT | 2 | MDXR.PHA- | 2 | 0V | 2 | PIN.GND | PIN TO PIN | 2 | 0V | 0V
3 | MDZT.PHA+ | 3 | MDZT.PHB+ | 3 | LNDZT.OUT | 3 | LNDYT.OUT | 3 | MDZT.PHB+ | 3 | MDX1.PHB+ | 3 | 0V | 3 | 0V | 3 | MDXR.PHA+ | 3 | NC | 3 | PIN.-12VL |  | 3 | NC | NC
4 | MDZT.PHA- | 4 | MDZT.PHB- | 4 | LPDYT.OUT | 4 | LNDX1.OUT | 4 | MDZT.PHB- | 4 | MDX1.PHB- | 4 | 24V | 4 | 24V | 4 | MDXR.PHA+ | 4 | EDXR.Z- | 4 | PIN.GND |  | 4 | EAYT.Z- | EAZT.Z-
5 | MDZT.PHB+ | 5 | 0V | 5 | LNDYT.OUT | 5 | LNDX2.OUT | 5 | MDYT.PHA+ | 5 | MDX2.PHA+ | 5 | LNDZT.OUT | 5 | LNDX1.OUT | 5 | MDXR.PHB- | 5 | EDXR.B- | 5 | PIN.OUT |  | 5 | EAYT.B- | EAZT.B-
6 | MDZT.PHB- | 6 | MDYT.PHA+ | 6 | LPDX1.OUT | 6 | 0V | 6 | MDYT.PHA- | 6 | MDX2.PHA- | 6 | 0V | 6 | 0V | 6 | MDXR.PHB- | 6 | EDXR.A- | 6 | PIN.GND |  | 6 | EAYT.A- | EAZT.A-
7 | MDZT.PHB+ | 7 | MDYT.PHA- | 7 | LNDX1.OUT | 7 | EDXR.A+ | 7 | MDYT.PHB+ | 7 | MDX2.PHB+ | 7 | 24V | 7 | 24V | 7 | MDXR.PHB+ | 7 | 5V | 7 | PIN.IN |  | 7 | 5V | 5V
8 | MDZT.PHB- | 8 | MDYT.PHB+ | 8 | LPDX2.OUT | 8 | EDXR.B+ | 8 | MDYT.PHB- | 8 | MDX2.PHB- | 8 | LPDYT.OUT | 8 | LPDX2.OUT | 8 | MDXR.PHB+ | 8 | 5V | 8 | PIN.CAL |  | 8 | 5V | 5V
9 | 0V | 9 | MDYT.PHB- | 9 | LPDX2.OUT | 9 | EDXR.Z+ |  |  |  |  | 9 | 0V | 9 | 0V | 9 | MDXR.LP | 9 | 0V | 9 | PIN.GAIN |  | 9 | 0V | 0V
10 | MDYT.PHA+ | 10 | MDX1.PHA+ | 10 | 5V | 10 | EDXR.P |  |  |  |  | 10 | 24V | 10 | 24V | 10 | MDXR.LN | 10 | EDXR.Q |  |  |  | 10 | EAYT.Q | EAZT.Q
11 | MDYT.PHA- | 11 | MDX1.PHB+ | 11 | 5V | 11 | EDXR.Q |  |  |  |  | 11 | LNDYT.OUT | 11 | LNDX2.OUT | 11 | NC | 11 | EDXR.P |  |  |  | 11 | EAYT.P | EAZT.P
12 | MDYT.PHA+ | 12 | MDX2.PHA+ | 12 | EDXR.A+ | 12 | 0V |  |  |  |  | 12 | 0V | 12 | 0V | 12 | NC | 12 | EDXR.Z+ |  |  |  | 12 | EAYT.Z+ | EAZT.Z+
13 | MDYT.PHA- | 13 | MDX2.PHB+ | 13 | EDXR.A- | 13 | MDXR.LP | [8 empty] | 13 | NC | 13 | EDXR.B+ |  |  |  | 13 | EAYT.B+ | EAZT.B+
14 | MDYT.PHB+ | 14 | 0V | 14 | EDXR.B+ | 14 | LPDZT.OUT | [8 empty] | 14 | 0V | 14 | EDXR.A+ |  |  |  | 14 | EAYT.A+ | EAZT.A+
15 | MDYT.PHB- | 15 | MDXR.PHA+ | 15 | EDXR.B- | 15 | LPDYT.OUT | [8 empty] | 15 | NC | 15 | NC |  |  |  | 15 | NC | NC
16 | MDYT.PHB+ | 16 | MDXR.PHA- | 16 | EDXR.Z+ | 16 | LPDX1.OUT
17 | MDYT.PHB- | 17 | MDXR.PHB+ | 17 | EDXR.Z- | 17 | LPDX2.OUT
18 | 0V | 18 | MDXR.PHB- | 18 | 0V | 18 | 5V
19 | MDXR.PHA+ | 19 | 0V | 19 | EDXR.P | 19 | 5V
20 | MDXR.PHA- | 20 | MDZT.PHA+ | 20 | EDXR.Q | 20 | EDXR.A-
21 | MDXR.PHA+ | 21 | MDZT.PHA- | 21 | 0V | 21 | EDXR.B-
22 | MDXR.PHA- | 22 | MDZT.PHB+ | 22 | MDXR.LP | 22 | EDXR.Z-
23 | MDXR.PHB+ | 23 | MDZT.PHB- | 23 | MDXR.LN | 23 | 0V
24 | MDXR.PHB- | 24 | 0V | 24 | 0V | 24 | 0V
25 | MDXR.PHB+ | 25 | MDYT.PHA+ | 25 | 0V | 25 | MDXR.LN
26 | MDXR.PHB- | 26 | MDYT.PHA-
27 | 0V | 27 | MDYT.PHB+
28 | MDX1.PHA+ | 28 | MDYT.PHB-
29 | MDX1.PHA- | 29 | MDX1.PHA-
30 | MDX1.PHB+ | 30 | MDX1.PHB-
31 | MDX1.PHB- | 31 | MDX2.PHA-
32 | 0V | 32 | MDX2.PHB-
33 | MDX2.PHA+ | 33 | 0V
34 | MDX2.PHA- | 34 | MDXR.PHA+
35 | MDX2.PHB+ | 35 | MDXR.PHA-
36 | MDX2.PHB- | 36 | MDXR.PHB+
37 | 0V | 37 | MDXR.PHB-

## SLITS SIGNALS
| DB25-1 | SLMOT |  | DB25-2 | SLENC |  | HUBER |  | PRECISTEP_1 |  | PRECISTEP_2 |  | LIMIT_SWITCH_1 |  | LIMIT_SWITCH_2
1 | MSX1.PHA+ | MSX1.PHA+ | 1 | 24V | 24V | 1 | MSXR.PHA- | 1 | MSX1.PHA+ | 1 | MSZ1.PHA+ | 1 | 24V | 1 | 24V
2 | MSX1.PHA- | MSX1.PHB+ | 2 | LPSX1.OUT | LNSX1.OUT | 2 | MSXR.PHA- | 2 | MSX1.PHA- | 2 | MSZ1.PHA- | 2 | LPSX1.OUT | 2 | LPSZ1.OUT
3 | MSX1.PHB+ | 0V | 3 | LNSX1.OUT | LNSX2.OUT | 3 | MSXR.PHA+ | 3 | MSX1.PHB+ | 3 | MSZ1.PHB+ | 3 | 0V | 3 | 0V
4 | MSX1.PHB- | MSX2.PHA+ | 4 | LPSX2.OUT | 0V | 4 | MSXR.PHA+ | 4 | MSX1.PHB- | 4 | MSZ1.PHB- | 4 | 24V | 4 | 24V
5 | MSX2.PHA+ | MSX2.PHB+ | 5 | LNSX2.OUT | LNSZ1.OUT | 5 | MSXR.PHB- | 5 | MSX2.PHA+ | 5 | MSZ2.PHA+ | 5 | LNSX1.OUT | 5 | LNSZ1.OUT
6 | MSX2.PHA- | 0V | 6 | LPSZ1.OUT | LNSZ2.OUT | 6 | MSXR.PHB- | 6 | MSX2.PHA- | 6 | MSZ2.PHA- | 6 | 0V | 6 | 0V
7 | MSX2.PHB+ | MSZ1.PHA+ | 7 | LNSZ1.OUT | 0V | 7 | MSXR.PHB+ | 7 | MSX2.PHB+ | 7 | MSZ2.PHB+ | 7 | 24V | 7 | 24V
8 | MSX2.PHB- | MSZ1.PHB+ | 8 | LPSZ2.OUT | MSXR.O2 | 8 | MSXR.PHB+ | 8 | MSX2.PHB- | 8 | MSZ2.PHB- | 8 | LPSX2.OUT | 8 | LPSZ2.OUT
9 | MSZ1.PHA+ | 0V | 9 | LNSZ2.OUT | MSXR.LP | 9 | MSXR.LP |  |  |  |  | 9 | 0V | 9 | 0V
10 | MSZ1.PHA- | MSZ2.PHA+ | 10 | 0V | 0V | 10 | MSXR.LN |  |  |  |  | 10 | 24V | 10 | 24V
11 | MSZ1.PHB+ | MSZ2.PHB+ | 11 | 12V | 12V | 11 | 12V |  |  |  |  | 11 | LNSX2.OUT | 11 | LNSZ2.OUT
12 | MSZ1.PHB- | 0V | 12 | MSXR.LP | 0V | 12 | MSXR.O2 |  |  |  |  | 12 | 0V | 12 | 0V
13 | MSZ2.PHA+ | MSXR.PHB+ | 13 | MSXR.LN | NC | 13 | MSXR.O1
14 | MSZ2.PHA- | MSXR.PHB- | 14 | MSXR.O1 | LPSX1.OUT | 14 | 0V
15 | MSZ2.PHB+ | 0V | 15 | MSXR.O2 | LPSX2.OUT | 15 | NC
16 | MSZ2.PHB- | MSXR.PHA+ | 16 | 0V | 0V
17 | MSXR.PHA+ | MSXR.PHA- | 17 |  | LPSZ1.OUT
18 | MSXR.PHA- | 0V | 18 |  | LPSZ2.OUT
19 | MSXR.PHA+ | 0V | 19 |  | 0V
20 | MSXR.PHA- | MSX1.PHA- | 20 |  | MSXR.O1
21 | MSXR.PHB+ | MSX1.PHB- | 21 |  | MSXR.LN
22 | MSXR.PHB- | 0V | 22 |  | 0V
23 | MSXR.PHB+ | MSX2.PHA- | 23 |  | 12V
24 | MSXR.PHB- | MSX2.PHB- | 24 |  | 0V
25 | 0V | 0V | 25 |  | NC
26 |  | MSZ1.PHA-
27 |  | MSZ1.PHB-
28 |  | 0V
29 |  | MSZ2.PHA-
30 |  | MSZ2.PHB-
31 |  | 0V
32 |  | MSXR.PHB+
33 |  | MSXR.PHB-
34 |  | 0V
35 |  | MSXR.PHA+
36 |  | MSXR.PHA-
37 |  | 0V